## Slide 1
[ Tile light yield and uniformity ]
Miguel Bernardo, João Gentil, Ricardo Gonçalo
FCC-hh Calo Meeting – 6 March 2018
Notes: 1

## Slide 2: Summary
Our goal is to characterize real-sized plastic tiles as planned for the FCC HCAL
Measure of the light yield and uniformity
Use realistic readout:
      Wavelength shifting optical fibre of appropriate length
      Readout by PMT – no SiPM available at the moment
      Outline:
      Cutting tiles to realistic sizes
      Experimental setup
      Light yield measurements

## Slide 3: Scintillating Tiles
Used one #7 tile from ATLAS TileCal pre-production to obtain:
One half-tile from layer 1
One tile from layer 1
One half-tile from layer 10
ATLAS
Tile #7

## Slide 4: Scintillating Tiles
Used one #7 tile from ATLAS TileCal pre-production to obtain:
One half-tile from layer 1
One tile from layer 1
One half-tile from layer 10
Tiles wrapped in TYVEK to increase uniformity

## Slide 5: Experimental Setup
Source mounted on a X-Y arm
90Sr/90Y β- source
Emax = 546 keV / 2.283 MeV, A ≈ 0.5 mCi
Optical readout: wavelength-shifting fibre
Readout: photomultiplier in current mode
Light yield is with respect to reference plastic scintillator tile
Photomultiplier tube
WLS fibre reading reference tile
WLS fibre reading test tile
Reference scintillating tile
Test scintillating tile

## Slide 6: Results so far
Tiles read by 2.3 m WLS fibre
Tiles NOT wrapped in TYVEK
WLS fibre contact-coupled to tile
Each point corresponds to 30 current measurements
Error bar is standard deviation
NOTE: Initial measurements with lots of electronic noise (especially when measuring small tile)
Response uniformity of tile #1 (note relative yield)
Response = 100 x (Signal – Noise) / Reference
X – 400 steps / cm

## Slide 7: Results so far
Tiles read by 2.3 m WLS fibre
Tiles wrapped in TYVEK
WLS fibre contact-coupled to tile
Held in place by TYVEK envelope
Each point corresponds to 30 current measurements
Error bar is standard deviation
Curves in figure:
Nearer to PMT
Middle
Far end
Response uniformity of tile #1 (note relative yield)
Response = 100 x (Signal – Noise) / Reference
X – 400 steps / cm

## Slide 8: Results so far
Response uniformity of tile #10 (note relative yield)

## Slide 9: Results so far
Response uniformity of tile #10 (note relative yield)
X (400 steps / cm)
X (400 steps / cm)

## Slide 10
Outlook
19
First measurements done but many still missing:
Measure tile #10 with shorter fibre
Comparing uniformity with/without TYVEK
Comparing with known TileCal tile – to normalize response
Ideally get hold of SiPM to repeat measurements
Notes: 10

## Slide 11
Bonus slides
Comments welcome!
Notes: 11

## Slide 12
288 mm
302 mm
100 mm
9 mm
72 mm
91.7°
91.7°
138 mm
144 mm
100 mm
Corte e polimento
69 mm
Telha tipo 7 de ATLAS TileCal
Hipótese B (Telha #2): obter um trapézio com altura de 100 mm de altura e bases de 144 e 138 mm, mais um meio trapézio cortado pelo eixo de simetria

## Slide 13
288 mm
302 mm
9 mm
250 mm
107 mm
113 mm
6 mm
Corte e polimento
Telha tipo 7 de ATLAS TileCal
Hipótese B (Telha #3): obter meio trapézio com altura de 250 mm e bases de 226/2 = 113 mm e 214/2 = 107 mm, cortado pelo eixo de simetria

## Slide 14: Experimental setup
6 – high-voltage source
7 – digital multimeter to read PMT current
8 radioactive source on X-Y arm

## Slide 15: Test using ATLAS #7 tile