## Slide 1: HGTD HV Patch Panels Status
Bengt Lund-Jensen, Lei Fan, Ricardo Gonçalo

## Slide 2: HGTD HV Patch Panels
Each of the 8032 HGTD modules will need a bias voltage between -300 V and -900 V, adjusted individually
Electronic noise induced in the DC bias voltage at the power supply or in the cables between USA15 and the detector must be filtered out.
Filter modules (EC-PP) will be installed on the end-cap calorimeter surfaces – filter noise and allow HV channel routing

## Slide 3: Patch Panel design parameters
Materials to withstand radiation and magnetic field:
  TID 15.0 Gy and 1 MeV neq fluence 1.0 × 1012 cm−2
  Magnetic field up to 0.5 Tesla
  Avoid easily activated and magnetic materials
  Avoid extensive use of dielectrics
Mechanical stability and ease of access during shutdowns
  Robust connectors
  Fixation to Tilecal and cable strain relief staves – to be studied together with Technical Coordination
Space constraints
  Around 20cm free in radial direction
  Number of wires should present good match to cables from HV supplies

## Slide 4: Patch Panel Filter Design
HV routing through fixed wire connections between input long cables from USA15 and short cables to detector modules
2nd order RC-RC low-pass filter to suppress AC noise
  Up to -900 V with, no significant leakage, supply currents up to 3 mA per channel
  Decoupling capacitors to suppress common-mode noise (C3, C4)

## Slide 5: Patch Panel Units Design
A modular design is proposed for the patch panels
Individual modules are aluminium boxes containing two filter boards and connectors
  Provide mechanical support and insulate each pair of boards within separate Faraday cage
  Easy to construct, handle and access for maintenance
  14 RC-RC low-pass filters in each filter board
  Means one 56-wire cable connected to each module: = 28 HV channels = 14 channels x 2 boards
  Routing of individual HV channels through wires connecting cables to each filter board

## Slide 6: Patch Panel design parameters
Differential HV channels, insulated from the patch panel unit
  Module aluminium boxes act as Faraday cage, electrically connected to the Tilecal surface
  HGTD earth extends through cable sleeve to filter board backplanes
  HGTD/Filter earth insulated from module earth (=Tilecal earth) to avoid ground loops

## Slide 7: Initial (pre-)prototype tests
RC–RC filter: 10 kΩ x 33 nF or 47 nF
Note R and C not final – must be defined together with sensors and tested together
Assuming 3 mA maximum current means:
60 V voltage drop
180 mW max. dissipation per channel
I.e. 5 W / 24-channel box
Ideal response:
fc = 338 Hz (33 nF) or 482 Hz (47 nF)
-40 dB / decade
Achieved this when tested inside shielding box
Chart: dB vs log10(f)
| 1.0 | 2.0 | 2.397940008672037 | 2.69897000433602 | 2.8750612633917 | 3.0 | 3.397940008672037 | 3.69897000433602 | 3.8750612633917 | 4.0
dB | -2.1026 | -2.2702 | -7.4322 | -13.4324 | -18.4891 | -22.4988 | -32.9563 | -39.1721 | -43.7417 | -46.0206
Chart: 
| 1.0 | 2.0 | 2.397940008672037 | 2.69897000433602 | 2.8750612633917 | 3.0 | 3.397940008672037 | 3.69897000433602 | 3.8750612633917 | 4.0
dB (0V) | -2.1026 | -2.2702 | -7.4322 | -13.4324 | -18.4891 | -22.4988 | -32.9563 | -39.1721 | -43.7417 | -46.0206
dB (1 kV) | -33.4324 | -33.8501 | -34.6331 | -40.9151 | -49.1186 | -53.9794 | -57.7211 | -73.9794 | -75.563

## Slide 8: Prototype tests and quality control
Planned tests:
  Connectivity of components and filter performance tested by measuring the filter response as function of frequency and load
  Leakage current
  Insulation between internal and external ground.
  Cross talk between channels
  Temperature under load
  Long term reliability including enhanced aging by temperature cycling in a climate chamber
  Radiation and magnetic field tolerances
For production:
  Based on prototype results will establish set of quality control benchmarks to be done in production and upon delivery

## Slide 9
Thanks!

## Slide 10: Prototype tests
Response improves with shielding (Faraday cage )
Closer to -40 dB / decade
Response flat above 10 kHz
Other tests foreseen:
Leak current under HV bias – waiting for precision HV module to be free, to ease measurement
Final tests must be done with HGTD HV source prototype
Chart: 
| 0.0 | 1.0 | 2.0 | 2.397940008672037 | 2.69897000433602 | 2.8750612633917 | 3.0 | 3.176091259055681 | 3.301029995663979 | 3.397940008672037 | 3.477121254719662 | 3.69897000433602 | 3.8750612633917 | 4.0 | 5.0 | 5.698970004336013 | 6.0
Shielding; direct; 33 nF | 0 | -0.33199 | -2.4988 | -7.2803 | -13.4324 | -18.0444 | -21.5403 | -27.5869 | -32.0412 | -36.1934 | -38.2373 | -46.4661 | -56.4782 | -58.0618 | -58.0618 | -58.0618 | -58.0618 | 0 | -0.67717 | -3.098 | -7.9588 | -13.7968 | -17.8898 | -21.6709 | -27.4322 | -32.9563 | -36.4782 | -38.9769 | -47.4322 | -55.1392 | -66.0206 | -66.0206 | -66.0206 | -66.0206
Shielding; direct; 47 nF | 0 | -0.67717 | -3.098 | -7.9588 | -13.7968 | -17.8898 | -21.6709 | -27.4322 | -32.9563 | -36.4782 | -38.9769 | -47.4322 | -55.1392 | -66.0206 | -66.0206 | -66.0206 | -66.0206
No shielding; direct; 33 nF | -2.1026 | -2.2702 | -7.4322 | -13.4324 | -18.4891 | -22.4988 | -32.9563 | -39.1721 | -43.7417 | -46.0206 | -46.0206 | -46.0206 | -46.0206

## Slide 11: Response curve
Response after adding common-mode decoupling capacitors

## Slide 12: Optimising design
Also tried an alternative design to:
Ease assembly: reduces production time and manufacture errors
Improve robustness: assess robustness of routing cables / soldering
Add decoupling capacitors against common-mode noise in each channel
Lower cost:
  Especially by reducing length of cable for HV routing
  Trying routing in 4-layer PCB

## Slide 13: Connectors
Found potentially interesting connectors from Farnell:
  http://www.farnell.com/datasheets/2916873.pdf
  Unit price (120 pins, small quant.): 53 € plug; 45 € pins; 26 € connector
  To be used for this prototype and replaced later
REF:516-120-000-101 REF:516-056-000-301 REF:516-120-000-402 REF:516-056-000-402 REF:516-230-512 REF:516 230-556

## Slide 14: Measurements
Difficult to measure low-frequency behaviour with our current setup due to output capacitor in waveform generator