## Slide 1: HGTD High Voltage Status
Ricardo Gonçalo (LIP/Univ. Coimbra) for the HV PS and PP teams
HGTD Week at CERN, February 2023

## Slide 2: High-Voltage Power Supplies
Power supplies undergoing ageing tests at the factory
To be finished by the end of February
After that will proceed with power supply tests

## Slide 3: High voltage Patch Panels
HV filtered and routed between input long cables from USA15 and short cables to detector
  2nd order RC-RC low-pass filter to suppress AC noise
  8032 HV channels: at least 574 boards of 14 filters each
MSc student (Antonio Caramelo) setting up automated testing station for QC
  Earlier prototype being used for developing tests
New prototype going for production soon – should be ready for tests in few weeks

## Slide 4: Patch Panel Tests
Using usb oscilloscope to generate and acquire waveform and attenuation curve
Extract cutoff frequency (-3 dB) to use as measure of filter quality
  Expect f3db=126,4 Hz and -40dB/decade for this example
  Get 122.4±4 and investigating extra peak at the moment

## Slide 5: Next Steps
HV Power Supplies:
  Ageing tests in February followed by testing of full HV supply units and noise tests
  Finalize HV supply specifications in June
Patch panels:
  New (hopefully final) prototype next month: 1 month delay but no technical problems
  Further tests: leakage current, cross talk, rad. hardness and magnetic field, ageing
  QC progressing
High Voltage PDR passed in August 2022:
  https://indico.cern.ch/event/1190013/
  https://edms.cern.ch/ui/file/2775726/1/HGTD-HVPS-PDR_docx_cpdf.pdf
HV supply FDR meeting on 7 July 2023; PP FDR on 29 August (merge?)

## Slide 6: Bonus slides

## Slide 8: Patch Panel Filter Design For PDR
HV routing through fixed wire connections between input long cables from USA15 and short cables to detector modules
2nd order RC-RC low-pass filter to suppress AC noise
  Up to -900 V with, no significant leakage, supply currents up to 3 mA per channel
  Decoupling capacitors to suppress common-mode noise (C3, C4)

## Slide 9: Patch Panel Grounding & Shielding
Use filter module boxes as continuation of vessel Faraday cage
  Also DC connecttion of filter boards backplane to cage
DC connection of Faraday cage to local TileCal earth
  DC current not as bad as HF AC noise
Move common-mode filter capacitors to filter input instead of output
  Filter common mode current at input

## Slide 10: Prototype tests and quality control
Planned tests:
  Connectivity of components and filter performance tested by measuring the filter response as function of frequency and load
  Leakage current
  Insulation between internal and external ground.
  Cross talk between channels
  Temperature under load
  Long term reliability including enhanced aging by temperature cycling in a climate chamber
  Radiation and magnetic field tolerances
  NEW: test with and without DC ground
For production:
  Based on prototype results will establish set of quality control benchmarks to be done in production and upon delivery

## Slide 11: HGTD HV Patch Panels
Each of the 8032 HGTD modules will need a bias voltage between -300 V and -900 V, adjusted individually
Electronic noise induced in the DC bias voltage at the power supply or in the cables between USA15 and the detector must be filtered out.
Filter modules (EC-PP) will be installed on the end-cap calorimeter surfaces – filter noise and allow HV channel routing

## Slide 12: Patch Panel design parameters
Materials to withstand radiation and magnetic field:
  TID 15.0 Gy and 1 MeV neq fluence 1.0 × 1012 cm−2
  Magnetic field up to 0.5 Tesla
  Avoid easily activated and magnetic materials
  Avoid extensive use of dielectrics
Mechanical stability and ease of access during shutdowns
  Robust connectors
  Fixation to Tilecal and cable strain relief staves – to be studied together with Technical Coordination
Space constraints
  Around 20cm free in radial direction
  Number of wires should present good match to cables from HV supplies

## Slide 13: Patch Panel Units Design
A modular design is proposed for the patch panels
Individual modules are aluminium boxes containing two filter boards and connectors
  Provide mechanical support and insulate each pair of boards within separate Faraday cage
  Easy to construct, handle and access for maintenance
  14 RC-RC low-pass filters in each filter board
  Means one 56-wire cable connected to each module: = 28 HV channels = 14 channels x 2 boards
  Routing of individual HV channels through wires connecting cables to each filter board